Funnel Demo

This demo shows a model of Ion Funnel interface. It is not optimised, as you will see from simulations.

Figure above describes a model where ions from ion funnel connected through an octopole ion guide to a 3D trap.

Example 1. Ion Funnel Demo.

- Start SIMAX. Using button  load a model Funnel_Octopole_3Dtrap.wb8  from simulation folder.

- Press  button to fly ions. You will see how ions of several masses (100 ions in each group) start from the middle of Ion Funnel and transmitted into a 3D trap.

As you can see at present power supply settings 50Da ions are completely lost, 10% of 150Da ions are trapped at RF traps in the exit of ion funnel and very few ions are trapped in a 3D trap. This is not suprising because trapping efficiency of a 3D trap with direct injection is only few percent and with big mass discrimination. The other desing corrects this problem to some extend.

Example 2. Collecting ions in the back of octopole.

In this simulation there is an additional field in the exit of Octopole. This field is a uniform potential distribution which can be realised in practice by a resistive coating of the part of octopole rods. With a potential of 5V this field produces a DC potential well at the back of an octopole, where ions are collected after some flight time.

- Start SIMAX. Using button  load a model ResistiveOctopole_Collecting ions in the octopole.wb8  from simulation folder.

- Press  button to fly ions. Ions of several masses here start in the middle of octopole and after some time you will see how ions are collected at the back of the Octopole.

- Go to Phase Space of Fly page and press  button, then select for Y axis of the plot. You will see the following picture of the phase space during ion fly

- Go to Splat Statistics of Fly page and press  button on the top plot, then select  in order to count remaining ions in this region.  You will see something like this

- Now click Ions item in the tree view of the model. In the panel at the top of model tree view press   button and save ions in a file “Trapped in the back of Octopole_10ms.ic8” . You can use this group in further simulations.

Example 3. Pulsed injection into a 3D trap.

After 10ms collection period ions are injected into a 3D trap by switching offsets of nearby electrodes. Before ion entrance into the trap the RF field inside the 3D trap is zero. It is switched on when majority of ions arrive into the trap. This method of ion trapping in a 3D trap is called “Fast RF start-up”. It was invented in Shimadzu Research Laboratories Ltd, Manchester, UK and is used by Shimadzu Co in commercial instruments. With appropriate optimisation it is possible to achieve 100% trapping of ions in a wide mass range.

- Start SIMAX. Using button  load a model ResistiveOctopole_Pulsed Injection into 3Dtrap.wb8  from simulation folder.

- Press  button to fly ions. You will see how ions from previous example start at the back of octopole and directed into the trap. The RF field inside the trap starts at 10.02ms and ions are trapped. Over 60% of ions are still lost.

- Go to Splat Statistics of Fly page and press  button on the top plot, then select  in order to count remaining ions in this region.  You will see something like this